## Slide 1
JEY Labs Presents
Test Gear
For
Hams
Doug Millar
K6JEY

## Slide 2
Doug Millar K6JEY
Modern Choices on a Budget

## Slide 3: Planning
Which gear do you use the most?
What will you need in the future?
Will you want to keep or get any favorites?

## Slide 4: MULTIMETERS

## Slide 5: They all read correctly with 50ohms
Fluke
8060a
(1977)
87iii
(2000)
Fluke 77
(1980)
Mastech 8906
(1999)
Fluke 115
(2010)

## Slide 6: You Can Keep an old Friend
But-
You get Tired of-
Too many
leads
switches
scales
Zeroings
Warm up time
It is still a good
RF voltmeter.

## Slide 7: Fluke 115 and 77 Zeros fast, auto ranging and tough
True RMS
One handed
About $110
$50

## Slide 8: ANENG 681
Color model available
Form factor of phone
USB charging
Excellent display
Auto function
Well built and sturdy
$26 Best Bang/$

## Slide 9: UT 181A
DMM with Android APP
via Bluetooth..
Does remote data
collection.
About $290
Also a cheaper model
UT61E+ at $90
But no internal data logging
Or DBm measurement.

## Slide 10: Meet Pencil DMM
Turn the phone
clockwise for this display
Turn it anticlockwise
and get the graphing
display.
You can take pictures of
the measurements.
No test lead
wires!
Did I say $35 from
Amazon?

## Slide 11: Fluke 8840A with 4 wire ohms Cheap, stable and reads to .001ohm!

## Slide 12: 4 Wire Ohms Measurement
The technique extends
The internal circuit to
The outside.
Mainly for measurements below 1K ohms

## Slide 13: AC/DC clamp on ammeter UT210E
About
$45
Reads
Down to
1ma DC

## Slide 14: Power Meter

## Slide 15: Power Meter Bird 4410A
Seven ranges
per element
5% of reading!
$300 with
elements

## Slide 16: Daiwa CN801hp
$60

## Slide 17: Newer Imported type.
$50
Has battery
12 watts
And other
Ranges.

## Slide 18: New Kid on the Block N-Type RF Power Meter V7 10GHz
0-10Ghz
-60dbm- 1watt
Offset
Charges
through C
$85

## Slide 19: Frequency Sources

## Slide 20
GPS oscillator
10MHz output
Data output
$150

## Slide 21: GPS Source with display
$190
Display shows
Actual Freq.

## Slide 22: Signal Generators

## Slide 23: Elecraft XG3 Signal Generator
$260
Accurate versatile
Ham band oriented
Generator.

## Slide 24: HP8642A/B DC to 2GHz generator
$500
Nothing newer beats it for frequency and attenuation
(Be sure and change the fan!).

## Slide 25: Agilent 8647/48 A-D

## Slide 26: WB-SG1 $168 You will also need an external attenuator $50-100

## Slide 27: External Attenuators
Narda
Step Atten.
Narda
Variable
Atten.

## Slide 28: Spectrum Analyzers

## Slide 29: Tiny SA Ultra
Has 200hz
Resolution BW
Will go to 10GHz

## Slide 30: 3KHz vs 300Hz RBW
KNX at 3KHz RBW
KNX at 300Hz RBW

## Slide 31: Block Diagram
1ST IF is 980MHz
2nd IF is 870KHz/SDR

## Slide 32: Tiny SA Ultra Inside
Attenuator, LNA, Input
Generator output
Mixer, filters
Switched filter

## Slide 33: Anritsu MT8212B
$400-800
Best all in one deal-

## Slide 34: MT8212b Screen

## Slide 35: Anritsu MS2712E 9Kc to 4GHz
$1500
Battery operated, hand held, GPS frequency correction.
$1500

## Slide 36: Anritsu m2712E Screen

## Slide 37: Oscilloscope
You want one with auto set
Analogue like display- 256 levels of intensity

## Slide 38: Excellent 100MHz scope <$100

## Slide 39: Rigol DHO802
70MHz- $299 USB c
Touchscreen, HDMI, 2” Thick!

## Slide 40: Counter

## Slide 41
Frequency
Counter
$115
13 digits
0-6GHz

## Slide 42: Miscellaneous

## Slide 43: LCR Meter DE 5000
LCR meter
ESR measurement
.01 ohm resolution
Less than $80

## Slide 44: MESR 100 v2 ESR Meter
Very good quality
It tells you what to do.
About $50

## Slide 45: SV34401 VNA
50KHz-4.4GHz
100db Dynamic Range
Full featured
Does TDR
$300

## Slide 46: NanoVNA-F Double check the calibration load
$120

## Slide 47: Resistance vs. Return Loss
Resistance SWR Return Loss
50.5 ohms 1.01 46db
51.0 1.02 40
51.5 1.03 36
52.0 1.04 34
52.5 1.05 32

## Slide 48
50.41
50.017 (M1408)
49.89
49.5
50.22
49.85
50.22
51.16
50.12
50.67
48.19
49.87
50.35
50.20
50.15
All below are 26.5GHz+
50.058 HP909d
50.080 Anritsu 40Ghz
50.040 Anritsu 40GHz
50.070 Winchel 40GHz
50.012 Anritsu 28V50 65GHz
50.264 HP85138 50GHz
Resistance Measurements of a Group Of 21 SMA 50 Ohm Loads

## Slide 49
Notice the difference in Strain relief
Cable shield and dielectric
To connector interface is the
problem.. Strain releif
Utiflex

## Slide 50: VNA Resources
From SDR-Kits $35
User’s Manual
HP 8720B VNA
(Download a PDF)

## Slide 51: Calibrators

## Slide 52: RCL Box
Handy calibration tool and cheap
Resistance
Capacitance
Inductance
EBay- “High Precision
Inductance Resistor Capacitor
LRC Calibrate Reference Module Box”
. Calibration to 5 digits. or so.
About $25

## Slide 53: RCL Box Cal Data
617.2uh @10KHz
623.0uh@100KHZ
5.7980 Ohms
19.819nf@10KHz
plus cal. date

## Slide 54: Find yourself a resistance standard
Check for
ESI SR-1
That is labeled
Honeywell

## Slide 55: Make your Own
You need a good resistor and a two pin banana plug
.005% at 10K ohm (VR marking)
Vishay RNC90Y10K000VR
$10 on EBay

## Slide 56: Calibration Data

## Slide 57: Voltage Source
Lab quality
Comes with cal. data
at $160
From
“Voltagestandard.com”

## Slide 58: Cheaper- 4 Output
$60 with PS
Has Cal Data (6 digits)
Very stable
LM399.
Well designed
Mine- 10.233214 v
Listing-
LM399 Power
Reference Module

## Slide 59: Good Design
Regulator is
insulated
Isolated
Power supply well
regulated
Good Temp co
resistors

## Slide 60: Save your eyes, do better work Koolertron 4.3” screen $70

## Slide 61: Contact Info
drzarkof56@yahoo.com

## Slide 62: That’s it. Any questions?